## Slide 1: Least square calibration
Bo Svensmark
Analytical Chemistry
PLEN
University of Copenhagen
svensmark@plen.ku.dk

## Slide 2: Linear calibration
Measure the response for a series of known concentrations
NO – this may not be OK at all!
Errors could be > 100 %
Plot the response as function of concentration
Make linear regression Y = a0 + a1X
Check the linear range
If R2 > 0.99 we conclude that: - the response is linear - we can use the calibration line for quantitative analysis

## Slide 3: What may be wrong with R2 >0.99
The random errors are not independent of concentration
In most cases random errors will be a combination of
Constant error – independent of concentration
Counting error – proportional to square root of concentration
Proportional error – proportional to concentration
What to do about random errors dependency of concentration:
Use weighted least square regression
Weights inverse proportional to the error2
Std. Dev.
Proportional
Square root
(most detectors)
Constant
Concentration

## Slide 4
To obtain the line: y = a0+ a1x
We must minimize the sum of the least squares (weighted):
xA
xB
Chart: 
| 4.3634134752287997E-4 | 2.2159242059690059E-3 | 8.7641502467842806E-3 | 2.6995483256594024E-2 | 6.4758797832945983E-2 | 0.12098536225957167 | 0.17603266338214973 | 0.19947114020071632 | 0.17603266338214973 | 0.12098536225957167 | 6.4758797832945983E-2 | 2.6995483256594024E-2 | 8.7641502467842806E-3 | 2.2159242059690059E-3 | 4.3634134752287997E-4
Y | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
A
B
Chart: 
| 4.3634134752287997E-4 | 2.2159242059690055E-3 | 8.7641502467842788E-3 | 2.6995483256594024E-2 | 6.4758797832945955E-2 | 0.12098536225957167 | 0.17603266338214973 | 0.19947114020071632 | 0.17603266338214973 | 0.12098536225957167 | 6.4758797832945955E-2 | 2.6995483256594024E-2 | 8.7641502467842788E-3 | 2.2159242059690055E-3 | 4.3634134752287997E-4
Y | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17
Least square linear regression
Error | Weight
Constant | C0
Square root | C½/X
Proportional | C1/X2
All (const + sqr + prop) | 1/(C0 + C½Sq(X) + C1X)2

## Slide 5: Input of weights in spreadsheet Calibration
The contributions is expressed as the standard deviation in the response at the conc. of the lowest standard in the calibration:
So: Constant uncertainty: sY independent of conc.
S½: Counting uncertainty: sY proportional to sqrt(conc.)
S1: Proportional uncertainty: sY proportional ton conc.
Std. Dev.
Proportional
Square root
(most detectors)
Constant
Concentration

## Slide 6: What may be wrong with R2 >0.99
2) There are systematic errors
The response may bend downward at high concentrations
The response may bend upward at low concentrations
What to do about systematic errors :
Use polynomial least square regression of degree > 1
Most often a quadratic equation: Y = a0 + a1X + a2X2

## Slide 7: What may be wrong with R2 >0.99
3) The concentrations in the calibration are not equidistant
This is normal for methods with MS-detection
Only a problem if:
  Uncertainty depends on concentration and/or
  There are systematic errors
THIS means: IT IS ALWAYS A PROBLEM for non-equidistant concentrations
Used for methods or problems with restricted dynamic range, e.g.:
AAS
UV/VIS spectrometry
Used for methods with long dynamic range, e.g.:
MS-detectors
FID in GC
Ionselective electrodes

## Slide 8: So –what to do?
Equidistant concentrations | Systematic error absent | Uncertainty independent of concentration | Unweighted linear | Weighted linear | Unweighted polynomial | Weighted polynomial
Yes | Yes | Yes | X |  |  | 
 |  | No |  | X |  | 
 | No | Yes |  |  | X | 
 |  | No |  |  |  | X
No | Yes | Yes | (x)* |  |  | 
 |  | No |  | X |  | 
 | No | Yes |  |  | (x)* | 
 |  | No |  |  |  | X
*Very unlikely
Linear or polynomial regression?
Compare R2 for degree 1, 2, …
This statistic test is very easy
Unweighted or weighted regression?
Look on the accuracy of the regression
Is the predicted concentrations with a good enough deviation from the correct value?
The Excel spreadsheet with all tests and calibration methods
is available at Absalon
Examples will be shown on the next slides

## Slide 9: Linear or polynomial regression?
Do it in Excel: Make linear regression with the Data Analysis -> Regression:
Look at the P-value for the highest degree of X:

## Slide 10: Unweighted or weighted regression?
Ok?

## Slide 11: Unweighted or weighted regression?
The accuracy of the regression can be estimated in this way:
Calculate the estimated X-values from the observed Y-values
Calculate the % error in the estimated X-values
Look at the error as function of X

## Slide 12: Unweighted or weighted regression?
So – even for R2 = 0.99994, the concentration inside the calibration range can be in error by 60 %

## Slide 13: Weighted linear regression, Wi = 1/Xi

## Slide 14: Repeat measurements
A) Dependent: Several injections from the same vial/solution
B) Independent: Several solutions prepared for each concentration or sample and injected from separate vials/solutions
Use the mean Y
N = no. of conc. = j
sY from i repeat injections
Xo
B) Use all Y as independent runs
N = repetitions*conc. = i*j
sY from i independent runs
Xo
Nominal conc.

## Slide 15: Repeat measurements
A) Dependent: Several injections from the same vial/solution
B) Independent: Several solutions prepared for each concentration or sample and injected from separatet vials/solutions
A) 3 injections of the same conc.
B) 3 injections of 3 slightly different conc.
Xo
Nominal conc.
Xo
B) But the results is assumed to be at the nominal conc.
Xo
Nominal conc.

## Slide 16: Errors in concentrations?
A) Dependent: Several injections from the same vial/solution
So, what is the best calibration function:
LINEAR with a bad fit to points
SECOND DEGREE with better fit?
All statistics points to 2nd degree
R2, Error %, Lack of Fit
But the most accurate would be:
LINEAR
Because it levels out the errors instead of modelling them

## Slide 17: Sequential dilutions
Start with the highest conc. and dilute to next conc., then dilute this to the next and so forth…
Y = 0.1*X One conc. error at X = 500

## Slide 18: Sequential dilutions
Start with the highest conc. and dilute to next conc., then dilute this to the next and so forth…
The problem is, that it LOOKS very nice even when it is wrong…

## Slide 19: Sequential dilutions
Start with the highest conc. and dilute to next conc., then dilute this to the next and so forth…
And furthermore, it looks as if the last point is in error – but it is actually the only correct point…

## Slide 20: Sequential dilutions
Y = 0.1*X 10 % error in dilutions up/down

## Slide 21: Sequential dilutions
Conclusion for sequential dilutions:
LOOKS NICE – AND HIDES ERRORS!

## Slide 22: Prepare calibration solutions independently
Errors will be at random concentrations giving:
MORE VARIATION – AND MORE CORRECT RESULTS
Be aware of SYSTEMATIC ERRORS (% dilutions ect.)
Be careful NOT TO OVERFIT (use too high degree and model noise)

## Slide 23: Conclusion
Always test R2 to see, if a linear or polynomial regression should be used
Always plot the data: –both for the whole range, and for low concentrations
Always look at the error in X to see if a weighted regression should be used
Never ever trust R2 alone as measure for linearity or goodness of fit
Do not overfit do not use too high order and model noise or error!
FINAL DECISION Look at the plots - use all your knowledge about the method!
Do not trust statistics blindly it is only a guide!
Do not use sequential dilutions looks nice, hides errors!

## Slide 24
Gannet - Sule